ICS 11.020

CCS C 10/29

T/CMEAS XXXX-XXXX

经颅电刺激技术在精神障碍临床应用中的操作规范

​Operation specifications for the use of transcranial electrical stimulation in mental disorders​

专家送审稿

xxxx - xx - xx发布

xxxx - xx - xx实施

中国医药教育协会发布

团体标准

目  次

前    言	II

1 范围	1

2 规范性引用文件	1

3 术语和定义	1

4 治疗场地	2

4.1 基本条件	2

4.2 电源要求	2

4.3 安全设施	2

5 设备	2

5.1 一般要求	2

5.2 电极要求	2

5.3 操作开始之前进行设备安全性检查	3

6 专业技术人员能力	3

6.1 经颅电刺激专业技术人员要求	3

6.2 上岗后教育	3

6.3 操作人员应有的急救能力	3

7 操作流程	3

7.1 经颅电刺激安全性筛查	4

7.2 患者准备	4

7.3 操作流程	4

8 tES诱发的不良反应、禁忌症和安全要求	6

8.1 通则	6

8.2 不良反应	6

8.3 禁忌症和安全建议	6

附录A（资料性）脑电图10-20定位系统	7

附录B（资料性）经颅电刺激推荐治疗方案表	8

参 考 文 献	11

前    言

本文件按本文件按照GB/T 1.1-2020《标准化工作导则 第1部分：标准化文件的结构和起草规则》的规定起草。

请注意本文件的某些内容可能涉及专利。本文件的发布机构不承担识别专利的责任。

本文件由浙江大学医学院附属第一医院提出。

本文件由中国医药教育协会归口。

本文件起草单位：浙江大学医学院附属第一医院、中南大学湘雅二医院、西安交通大学第一附属医院、武汉大学人民医院、浙江大学医学院附属邵逸夫医院、浙江大学医学院附属精神卫生中心/杭州市第七人民医院、上海市精神卫生中心、广州医科大学附属脑科医院、中国医科大学附属盛京医院、合肥合傲科技有限公司、江西杰联医疗设备有限公司、空军军医大学第一附属医院、南京医科大学附属脑科医院、广东省人民医院  四川大学华西医院、同济大学附属同济医院、复旦大学附属华山医院、上海交通大学医学院附属仁济医院  复旦大学附属中山医院  哈尔滨医科大学附属第一医院 昆明医科大学第一附属医院  天津市安定医院  山东第一医科大学附属省立医院  山西医科大学第一医院 郑州大学第一附属医院 北京大学第六医院、湖州市第三人民医院、宁波大学附属康宁医院、温州医科大学附属第一医院、浙江大学、南京沃高医疗科技有限公司、西安科悦医疗技术有限公司、西安穹顶医疗科技有限公司

本文件主要起草人：黄满丽、胡少华、王小平、张燕、马现仓、刘忠纯、陈炜、李涛、程宇琪、王振、袁逖飞、彭代辉、宁玉萍、汤艳清、张效初、李小俚、王化宁、姚志剑、何红波、王强、申远、吴珩、邵春红、骆艳丽、黄啸、夏炎、许秀峰、李申、王育梅、刘志芬、宋学勤、邓佳慧、沈鑫华、禹海航、杨闯、斯科、李上达、王芳、王聪、李龙

经颅电刺激技术在精神障碍临床应用中的操作规范

范围

本文件规定了经颅电刺激技术在精神障碍临床应用的场地、设备、专业技术人员的能力、操作流程等内容。

本文件适用于经颅电刺激技术在各类精神障碍中的临床应用。

规范性引用文件

下列文件中的内容通过文中的规范性引用而构成本文件必不可少的条款。其中，注日期的引用文件，仅该日期对应的版本适用于本文件；不注日期的引用文件，其最新版本（包括所有的修改单）适用于本文件。

GB 9706.1-2020 医用电气设备 第1部分：基本安全和基本性能的通用要求。

YY 9706.102-2021 医用电气设备 第1-2部分：基本安全和基本性能的通用要求 并列标准：电磁兼容 要求和试验

术语和定义

下列术语和定义适用于本文件。

  - 经颅直流电刺激 transcranial direct current stimulation；tDCS

通过在头皮表面放置电极，施加恒定强度(通常1-2mA)的直流电流，可调节大脑皮层的兴奋性。根据电流方向的不同，阳极刺激增加皮层兴奋性，阴极刺激降低皮层兴奋性。

  - 经颅交流电刺激 transcranial alternating current stimulation；tACS

通过在头皮表面放置电极，施加正弦、双相交流电流，电流强度和方向随时间周期性改变。

  - 经颅脉冲电刺激 transcranial pulsed current stimulation；tPCS

通过在头皮表面放置电极，施加单相或双相脉冲电流，进行刺激。

  - 经颅随机噪声刺激 transcranial random noise stimulation；tRNS

通过在头皮表面放置电极，施加随机振幅分布的电流，进行刺激。根据使用频率范围的不同，可分为全频谱噪声（通常0.1-640Hz）和高频噪声（通常100-640Hz）。

  - 经颅振荡直流电刺激 Oscillatory transcranial direct current stimulation；otDCS

通过在头皮表面放置电极，施加正弦波交流电流和直流电流，使电流以特定频率呈正弦波振荡，且始终保持单一极性，可以同时调节神经元膜电位和振荡节律。

  - 高精度经颅电刺激 high-Definition transcranial electrical stimulation；HD-tES

在头皮放置多个小尺寸电极（通常中心一个刺激电极，周围4-8个回路电极，电极直径约12mm）代替传统两片大尺寸海绵电极（25-35cm²），进行直流电刺激。

  - 时间干涉电刺激 Temporal Interference Stimulation；TIS

通过在头皮表面施加两个频率相近（如2000Hz和2010Hz）的高频正弦交流电流，形成低频（如10Hz）的包络调制，选择性地调节深部脑区的神经活动。

治疗场地

基本条件 ：

场地面积≥15㎡，温度18~26℃；湿度40~70%；照度300~750lux。

电源要求

经颅电刺激设备的功率为10W-65W，电源要求输入AC100-240V~50-60Hz，1A。

安全设施

配备灭火器等消防器材、抢救车等急救设备（包括药物）及血压计、体温计等、应急照明、通风设施。

设备

一般要求

经颅电刺激设备应获得有效的医疗器械注册证书。

设备应符合GB 9706.1-2020《医用电气设备 第1部分：基本安全和基本性能的通用要求》的规定。

设备应符合YY 9706.102-2021《医用电气设备 第1-2部分：基本安全和基本性能的通用要求 并列标准：电磁兼容 要求和试验》的规定。

电极要求

电极类型

常规经颅电刺激所用电极：由一阴一阳两个矩形海绵电极组成，电极的尺寸和电流强度共同决定是否能达到治疗所需的电流密度，常见的电极尺寸为5cm×5cm或5cm×7cm。

高精度经颅电刺激所用电极：直径通常为12mm的氯化银环形电极，要求具有较好的电流聚焦性和较高的空间分辨率。

电极管理

使用期限：一次性电极，单次使用；可重复使用电极，使用次数不超过30次；定期检查电极完整性。刺激参数要求

明确基本参数范围

电流强度：电流强度一般在0-2mA。TIS在0~5mA，高强度tACS可达15mA。

刺激频率：tACS、tPCS、tRNS频率满足在0~200Hz 范围内，除tDCS、tRNS外，tACS、tPCS、toDCS需连续可调，TIS频率范围满足0~2200Hz。

明确波形要求

tDCS：直流；tACS：正弦波；tPCS：方波；tRNS：随机噪声；tODCS：方波、正弦波、随机噪声；TIS：两组频率不同的正弦波。

明确功能模块要求

治疗功能：刺激、预刺激、伪刺激、一键终止、实时电流调节。

监测功能：实时阻抗监测、电流强度监测、治疗时间监测。

安全性检查

操作开始之前应进行设备安全性检查。

专业技术人员能力

经颅电刺激专业技术人员要求

应由经过执业医师、医技、护士担任。在接受经颅电刺激相关培训、进修，包括经颅电刺激的基本原理、一系列的操作规范培训，能熟练运行仪器后进行临床操作。

上岗后教育

应接受每年的继续教育和督导，进一步规范和提高标准化能力。

操作人员应有的急救能力

须要指出，尽管目前国际各类安全性指南认为，tES与癫痫发作不存在因果关系，但施加多次高强度的tACS电流，如15mA，仍然可能在干预后出现癫痫样放电，即使不诱发癫痫发作，操作人员仍应具有一定的急救能力，熟练掌握心肺复苏、癫痫发作的处理流程和方法。结合国际临床神经电生理联盟（International Federation of Clinical Neurophysiology，IFCN）2021年指南所推荐的癫痫发作应对流程及2025年中国抗癫痫协会发布的院前紧急终止癫痫发作处理原则，建议为：

- 保持冷静；

- 呼救；

- 确保环境安全，避免病人受到伤害（例如，帮助他们卧倒在地上并移走周边可能伤害他们的东西）；

- 记录癫痫发作的时间；

- 关注患者生命体征变化，直到救援到来；

- 适度宽松衣服；

- 保持呼吸道通畅，防止误吸（如将病人头偏向一侧）。此外，应注意避免试图约束病人或将东西放入病人的嘴里。

- 如持续5分钟以上或有其他严重症状，需尽快急诊就诊，患者照护者可考虑苯二氮卓类药物，经直肠、口腔黏膜、鼻粘膜等给药，用药后需注意呼吸变慢、血氧下降等情况。

操作流程

经颅电刺激安全性筛查

根据tES的禁忌症和不良反应，相关安全性指南中提到的一些关注点，制定的患者安全性筛查，有如下情况的，需谨慎使用经颅电刺激治疗：

- 头颅内或头皮上有金属物品或装置；

- 心脏起搏器；

- 颅骨孔洞或裂缝；

- 医学植入装置：心脏内导线、中心静脉导管；

- 皮质中风史或其他形式的脑器质性疾病，如脑瘤；

- 之前被诊断为抽搐或癫痫；

- 曾经做过神经外科手术；

- 已怀孕或可能怀孕；

- 偏头痛——是否可控？（药物控制或主观控制）；

- 身体有某些病症或不适尚未痊愈；

- 目前正在使用可降低抽搐发作阈值的药物，如三环类抗抑郁剂、心境稳定剂、锂盐、茶碱，或可能诱发出血的抗凝血剂药物。

患者准备

患者选择舒适的坐姿，在刺激之前，尽量保持头皮的清爽和干燥。

操作流程

常规双通道经颅电刺激

注：需确保海绵垫片与皮肤紧密接触，避免治疗过程中移位；需确保电极阻抗值＜50kΩ。如阻抗过高，需检查电极接触情况并调整。

高强度经颅电刺激

基于77.5Hz经颅交流电刺激提高脑脊液β‐内啡肽和5‐羟色胺的假说，高强度经颅电刺激技术的临床研究也逐步开展。流程类似于常规双通道经颅电刺激，常用方案为：电流为15mA，频率为77.5Hz， tACS 刺激20个疗程，每周治疗5次，每次40分钟，持续4周。

高精度五通道经颅电刺激

注：需确保高精度环形电极托安装牢固。进行阻抗检测时，需确保各通道阻抗值在设备要求的范围内（通常＜50 kΩ），如发现阻抗异常，及时检查并调整。

时间干涉电刺激

TI电极放置位置如图1所示，“O”为刺激靶点的位置，A1、A2、B1、B2为4个电极位置（这4个电极位置通过要刺激的靶点，计算获得）。A1-A2为一个通道，设置固定频率如2kHz，B1-B2为另一个通道，设置另一个固定频率，如2.01kHz，则两通道存在一个频率差，为10Hz。

在电极位置固定的情况下，可以通过调整电流比例来转移刺激焦点，如：通过改变两对电极之间的电流比例（保持总电流不变） 1:1比例：刺激点在中心；1:2.5比例：刺激点偏离中心20%；1:4比例：刺激点偏离中心35%。

需注意：确保导电膏均匀接触头皮，避免电流密度过高；避免使用表面氧化的电极。

图1

刺激靶点定位方法

规范的治疗参数需要配有相对规范的定位方法，目前，经颅电刺激定位最常用以下几种方式。

- 颅表标志点定位：

应用比较多的是国际脑电图10-20定位系统(见附录A)；

大脑皮质分成额叶、颞叶、顶叶、枕叶，经颅电刺激以这几个脑区为基础，采用国际脑电图10-20定位系统进行定位。脑电图电极放置方法来源于IFCN2017年更新版本，可根据该放置方法定位电极的刺激部位。

英文：根据脑解剖部位命名，额叶为F、颞叶为T、顶叶为P、枕叶为O。数字：左脑为奇数，右脑为偶数。

- 脑磁共振影像引导下精准定位。

基于磁共振(MRI)进行个体化精准定位经颅电刺激治疗是目前的研究趋势。认知任务fMRI，可以确定个体化的激活区，如手运动激活区、语言加工激活区等，可以根据激活的精准位置进行经颅电刺激治疗。

TI的MRI定位方法也是基于要刺激的深部脑区的MINI坐标和个体化的MRI数据，根据大脑的仿真结果，通过神经靶向软件，计算对该刺激靶区的最佳电极放置方案，以确定治疗靶点。

选择合适的治疗方案

详细的治疗方案见附录B。

根据现有的临床循证证据，本文件仅对tDCS和tACS的治疗方案进行推荐，其他电刺激包括经颅脉冲电刺激、经颅随机噪声刺激、经颅振荡直流电刺激、时间干涉电刺激的建议治疗方案等尚待更多的研究以明确。

tES诱发的不良反应、禁忌症和安全要求

通则

对tES诱发的不良反应进行知情告知。

不良反应

目前认为，主流方案（电流强度＜4mA）的tES，包括tDCS、tACS和tRNS均不会引起严重不良反应，常见的需要临床干预的不良反应包括：tDCS所致的皮肤灼伤（因电极与皮肤接触不佳所致）、抑郁症患者转躁（发生率0.06%左右）其他轻微的不良反应包括：电极处皮肤刺痛、轻微的疲劳、灼烧感、瘙痒感、头痛、电极处皮肤红斑等。大部分不良反应均为一过性，在刺激结束后可自行缓解。

至于严重不良反应癫痫，在6.3部分已经提到，目前临床研究、荟萃分析显示4mA以下的tES与癫痫发作不存在因果关系，也不会增加如脑卒中患者等共病躯体疾病患者的癫痫发作概率，但高强度的电流如15mA，仍然可能诱发癫痫样放电。

禁忌症和安全建议

对于tES禁忌症，使用安全方面的建议如下。

- 植入物或不可移除的颅内金属或装置不应使用经颅电刺激、服用已知可降低癫痫阈值药物的患者应谨慎使用经颅电刺激。

- 检查皮肤是否有任何预先存在的刺激、切口或病变——避免在受损皮肤和头骨病变上进行电刺激。

- 针对＞2岁的儿童及孕妇，使用经颅电刺激是安全的。

（资料性）

脑电图10-20定位系统

脑电图10-20定位系统见图A.1。

脑电图10-20定位系统

（资料性）

经颅电刺激推荐治疗方案表

根据《经颅直流电刺激治疗常见神经精神疾病的临床应用专家共识》（2022），介绍目前已有的各类精神障碍的经颅电刺激推荐治疗方案。

具体治疗参数根据目前国内外循证证据，推荐经颅电刺激临床应用方案，整理成表，如表B.1所示。

确定每个适应症的经颅电刺激有效性或无效性的证据水平，是通过应用于具有相同临床特征的患者的相同刺激方法获得的。对推荐等级做以下说明：

A级（疗效确切）采用窄置信区间或纳入重复双盲随机对照试验的 Meta 分析文献、至少2个 Ⅰ类研究、1个 Ⅰ类研究加至少2个Ⅱ类研究；

B级（疗效可信）采用宽置信区间的Meta分析文献、至少2个Ⅱ类研究、1个Ⅰ类研究加1个Ⅱ类研究、1个Ⅰ类研究加至少2个Ⅲ类研究、1个Ⅱ类研究加至少 2 个Ⅲ类研究；

C级（疗效可能）至少2个Ⅲ类研究、1个Ⅰ类研究加1个Ⅲ类研究、1个Ⅱ类研究加1个Ⅲ类研究。

D级（可期待）存在至少1个Ⅲ类研究、2个以上IV类研究，但证据尚不充分。

抑郁症推荐治疗方案

| 方案推荐 | 波形 | 阳极 | 阴极 | 电流密度（A/m²） | 刺激时间（min） | 治疗次数（次） | 证据等级 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 临床推荐1 | tDCS | F3 | F4 | ≥0.571 | ≥20 | 10~22 | A |
| 临床推荐2 | tDCS | F3 | 右侧眶上区 | ≥0.286 | ≥20 | 5~20 | A |
| 临床推荐3 | tDCS | F3 | F8 | ≥0.571 | ≥20 | 5~20 | B |

精神分裂症推荐治疗方案

| 方案推荐 | 波形 | 阳极 | 阴极 | 电流密度（A/m²） | 刺激时间（min） | 治疗次数（次） | 证据等级 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 临床推荐 | tDCS | F3和FP1中点 | T7和P3中点 | ≥0.515 | ≥20 | 10~40 | B |

物质使用所致障碍推荐治疗方案

| 方案推荐 | 波形 | 阳极 | 阴极 | 电流密度（A/m²） | 刺激时间（min） | 治疗次数（次） | 证据等级 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 尼古丁依赖临床推荐 | tDCS | F3 | F4 | ≥0.571 | ≥20 | 5~20 | A |
| 酒精依赖临床推荐 | tDCS | F4 | F3 | ≥0.571 | ≥13 | 1~10 | B |
| 甲基苯丙胺依赖临床推荐 | tDCS | F4 | F3或左侧眶上区 | ≥0.571 | ≥20 | 1~10 | C |

强迫症推荐治疗方案

| 方案推荐 | 波形 | 阳极 | 阴极 | 电流密度（A/m²） | 刺激时间（min） | 治疗次数（次） | 证据等级 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 临床推荐1 | tDCS | 右侧小脑 | Fp1 | ≥0.571 | 20 | 10 | C |
| 临床推荐2 | tDCS | Fz | 右侧眶上区 | ≥0.571 | 20 | 10~20 | C |

焦虑症推荐治疗方案

| 方案推荐 | 波形 | 阳极 | 阴极 | 电流密度（A/m²） | 刺激时间（min） | 治疗次数（次） | 证据等级 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 临床推荐 | tDCS | F3 | Fpz | ≥0.571 | 20 | 5~10 | B |

疼痛推荐治疗方案

| 方案推荐 | 波形 | 阳极 | 阴极 | 电流密度（A/m²） | 刺激时间（min） | 治疗次数（次） | 证据等级 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 神经性疼痛 | tDCS | 健侧C3/C4 | 患侧眶上 | 0.571 | 20 | 1~10 | B |
| 纤维肌痛 | tDCS | C3 | 右侧眶上区 | 0.571 | ≥20 | 5~10 | B |
| 偏头痛 | tDCS | C3 | 右侧眶上区 | ≥0.571 | 20 | 5~12 | B |

阿尔兹海默症推荐治疗方案

| 方案推荐 | 波形 | 阳极 | 阴极 | 电流密度（A/m²） | 刺激时间（min） | 治疗次数（次） | 证据等级 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 临床推荐 | tDCS | F3 | 右侧眶上区 | ≥0.571 | 20 | 5~10 | 可期待 |

轻度认知障碍症推荐治疗方案

| 方案推荐 | 波形 | 阳极 | 阴极 | 电流密度（A/m²） | 刺激时间（min） | 治疗次数（次） | 证据等级 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 临床推荐 | tDCS | F3 | 右侧眶上区 | ≥0.571 | 20 | 5~10 | 可期待 |

经颅交流电推荐治疗方案

| 波形 | 方案推荐 | 阳极 | 阴极 | 电流强度（mA） | 频率（Hz） | 刺激时间（min） | 治疗次数（次） | 证据等级 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| tACS | 抑郁症 | F3 | F4 | 1.5 | 40 | 20 | 10-20 | 可期待 |
| tACS | 精神分裂症（阴性症状） | F3 | F4 | 2 | 40 | 20 | 10-20 | 可期待 |
| tACS | 精神分裂症（幻听） | F3 | CP6 | 2 | 10 | 20 | 10-20 | 可期待 |
| tACS | 创伤后应激障碍 | 顶枕区域 | 对侧眶上区域 | 2 | 10 | 20 | 10-20 | 可期待 |
| tACS | 强迫症 | F3 | 右侧眶上区 | 2 | 40 | 20 | 10-20 | 可期待 |
| tACS | 阿尔兹海默症 | F3 | 右侧眶上区 | 1.5 | 40 | 20 | 10-20 | 可期待 |
| tACS | 轻度认知障碍 | F3 | 右侧眶上区 | 2 | 40 | 20 | 10-20 | 可期待 |

参 考 文 献

[1] Practical guide to transcranial direct current stimulation： principles, procedures and applications[M]. Springer, 2019.

[2] 唐睿,宋洪文,孔卓,等. 经颅直流电刺激治疗常见神经精神疾病的临床应用专家共识[J]. 中华精神科杂志,2022,55(5)：327-382.

[3] Russo C, Carneiro M I S, Bolognini N, et al. Safety review of transcranial direct current stimulation in stroke[J]. Neuromodulation： Technology at the Neural Interface, 2017, 20(3)： 215-222.

[4] Jaberzadeh S, Bastani A, Zoghi M. Anodal transcranial pulsed current stimulation： a novel technique to enhance corticospinal excitability[J]. Clinical Neurophysiology, 2014, 125(2)： 344-351.

[5] Murphy O W, Hoy K E, Wong D, et al. Transcranial random noise stimulation is more effective than transcranial direct current stimulation for enhancing working memory in healthy individuals： Behavioural and electrophysiological evidence[J]. Brain Stimulation, 2020, 13(5)： 1370-1380.

[6] Datta A, Bansal V, Diaz J, et al. Gyri-precise head model of transcranial direct current stimulation： improved spatial focality using a ring electrode versus conventional rectangular pad[J]. Brain stimulation, 2009, 2(4)： 201-207. e1.

[7] Kuo H I, Bikson M, Datta A, et al. Comparing cortical plasticity induced by conventional and high-definition 4× 1 ring tDCS： a neurophysiological study[J]. Brain stimulation, 2013, 6(4)： 644-648.

[8] Guo W, He Y, Zhang W, et al. A novel non-invasive brain stimulation technique：“Temporally interfering electrical stimulation”[J]. Frontiers in Neuroscience, 2023, 17： 1092539.

[9] Moffa AH, Martin D, Alonzo A, et al. Efficacy and acceptability of transcranial direct current stimulation (tDCS) for major depressive disorder： an individual patient data meta‐analysis[J]. ProgNeuropsychopharmacol Biol Psychiatry, 2020, 99：109836. DOI： 10.1016/j.pnpbp.2019.109836.

[10] Wang J, Luo H, Schülke R, et al. Is transcranial direct current stimulation, alone or in combination with antidepressant medications or psychotherapies, effective in treating major depressive disorder? A systematic review and meta‐analysis[J]. BMC Med, 2021, 19(1)：319.DOI： 10.1186/s12916‐021‐02181‐4.

[11] Razza LB, Palumbo P, Moffa AH, et al. A systematic review and meta‐analysis on the effects of transcranial direct current stimulation in depressive episodes[J]. Depress Anxiety, 2020, 37(7)：594‐608. DOI： 10.1002/da.23004.

[12] Kalu UG, Sexton CE, Loo CK, et al. Transcranial direct current stimulation in the treatment of major depression： a meta‐analysis[J]. Psychol Med, 2012, 42(9)：1791‐1800. DOI： 10.1017/S0033291711003059.

[13] Fregni F, Boggio PS, Nitsche MA, et al. Treatment of major depression with transcranial direct current stimulation[J]. Bipolar disorders, 2006, 8(2)： 203‐205.

[14] Fregni F, Boggio PS, Lima MC, et al. A sham‐controlled,

phase Ⅱ trial of transcranial direct current stimulation for the treatment of central pain in traumatic spinal cord injury[J]. Pain, 2006, 122(1‐2)： 197‐209. DOI： 10.1016/j.pain.2006.02.023.

[15] Lin YY, Chang CC, Huang CC, et al. Efficacy and neurophysiological predictors of treatment response of adjunct bifrontal transcranial direct current stimulation (tDCS) in treating unipolar and bipolar depression[J]. J Affect Disord, 2021, 280(Pt A)： 295‐304. DOI： 10.1016/j.jad.2020.11.030.

[16] Segrave RA, Arnold S, Hoy K, et al. Concurrent cognitive control training augments the antidepressant efficacy of tDCS： a pilot study[J]. Brain Stimul, 2014, 7(2)： 325‐331. DOI： 10.1016/j.brs.2013.12.008.

[17] Osoegawa C, Gomes JS, Grigolon RB, et al. Non‐invasive brain stimulation for negative symptoms in schizophrenia： an updated systematic review and meta‐analysis[J]. Schizophr Res, 2018, 197： 34‐44. DOI：10.1016/j.schres.2018.01.010.

[18] Cheng P, Louie L, Wong YL, et al. The effects of transcranial direct current stimulation (tDCS) on clinical symptoms in schizophrenia： a systematic review and meta‐analysis[J]. Asian J Psychiatr, 2020, 53：102392.

[19] Liu Y, Gu N, Cao X, et al. Effects of transcranial electrical stimulation on working memory in patients with schizophrenia： a systematic review and meta‐analysis[J].Psychiatry Res, 2021, 296： 113656. DOI：10.1016/j.psychres.2020.113656.

[20] Kantrowitz JT, Sehatpour P, Avissar M, et al. Significant improvement in treatment resistant auditory verbal hallucinations after 5 days of double‐blind, randomized,sham controlled, fronto‐temporal, transcranial direct current stimulation (tDCS)： a replication/extension study [J]. Brain Stimul, 2019, 12(4)： 981‐991. DOI： 10.1016/j.brs.2019.03.003.

[21] Boggio PS, Sultani N, Fecteau S, et al. Prefrontal cortex modulation using transcranial DC stimulation reduces alcohol craving： a double‐blind, sham‐controlled study[J].Drug and alcohol dependence, 2008, 92(1‐3)： 55‐60.

[22] Klauss J, Penido Pinheiro LC, Silva Merlo BL, et al. A randomized controlled trial of targeted prefrontal cortex modulation with tDCS in patients with alcohol dependence[J]. Int J Neuropsychopharmacol, 2014, 17(11)：1793‐1803. DOI： 10.1017/S1461145714000984.

[23] Klauss J, Anders QS, Felippe LV, et al. Multiple sessions of transcranial direct current stimulation (tDCS) reduced craving and relapses for alcohol use： a randomized placebo‐controlled trial in alcohol use disorder[J]. Front Pharmacol, 2018, 9：716. DOI： 10.3389/fphar.2018.00716.

[24] Klauss J, Anders QS, Felippe LV, et al. Lack of effects of extended sessions of transcranial direct current stimulation (tDCS) over dorsolateral prefrontal cortex on craving and relapses in crack‐cocaine users[J]. Front Pharmacol, 2018, 9：1198. DOI： 10.3389/fphar.2018.01198.

[25] Shahbabaie A, Golesorkhi M, Zamanian B, et al. State dependent effect of transcranial direct current stimulation (tDCS) on methamphetamine craving[J]. Int J Neuropsychopharmacol, 2014, 17(10)： 15911598.DOI：10.1017/S1461145714000686.

[26] Ghorbani Behnam S, Mousavi SA, Emamian MH. The effects of transcranial direct current stimulation compared to standard bupropion for the treatment of tobacco dependence： a randomized sham‐controlled trial[J]. Eur Psychiatry, 2019, 60： 41‐48. DOI： 10.1016/j.eurpsy.2019.04.010.

[27] Bation R, Mondino M, Le Camus F, et al. Transcranial direct current stimulation in patients with obsessive compulsive disorder： a randomized controlled trial[J]. Eur Psychiatry,2019, 62：38‐44. DOI： 10.1016/j.eurpsy.2019.08.011.

[28] Gowda SM, Narayanaswamy JC, Hazari N, et al. Efficacy of pre‐supplementary motor area transcranial direct current stimulation for treatment resistant obsessive compulsive disorder： a randomized, double blinded, sham controlled trial[J]. Brain Stimul, 2019, 12(4)： 922‐929.DOI： 10.1016/j.brs.2019.02.005.

[29] Jafari E, Alizadehgoradel J, Pourmohseni Koluri F, et al.Intensified electrical stimulation targeting lateral and medial prefrontal cortices for the treatment of social anxiety disorder： a randomized, double‐blind,parallel‐group, dose‐comparison study[J]. Brain Stimul,2021, 14(4)：974‐986. DOI： 10.1016/j.brs.2021.06.005.

[30] de Lima AL, Braga F, da Costa R, et al. Transcranial direct current stimulation for the treatment of generalized anxiety disorder： a randomized clinical trial[J]. J Affect Disord, 2019, 259：31‐37. DOI： 10.1016/j.jad.2019.08.020.

[31] Bae SH, Kim GD, Kim KY. Analgesic effect of transcranial direct current stimulation on central post‐stroke pain[J].Tohoku J Exp Med, 2014, 234(3)：189‐195. DOI： 10.1620/tjem.234.189.

[32] Attal N, Ayache SS, Ciampi De Andrade D, et al. Repetitive transcranial magnetic stimulation and transcranial direct‐current stimulation in neuropathic pain due to radiculopathy： a randomized sham‐controlled comparative study[J]. Pain, 2016, 157(6)：1224‐1231. DOI：10.1097/j.pain.0000000000000510.

[33] Fagerlund AJ, Hansen OA, Aslaksen PM. Transcranial direct current stimulation as a treatment for patients with fibromyalgia： a randomized controlled trial[J]. Pain, 2015,156(1)：62‐71. DOI： 10.1016/j.pain.0000000000000006.

[34] Dasilva AF, Mendonca ME, Zaghi S, et al. tDCS‐induced analgesia and electrical fields in pain‐related neural networks in chronic migraine[J]. Headache, 2012, 52(8)：1283‐1295. DOI： 10.1111/j.1526‐4610.2012.02141.x.

[35] Rasmussen et al. “High-definition tDCS improves delayed memory in Alzheimer’s disease patients： A pilot study using computational modeling to optimize electrode.” (2022).

[36] Zhang R, Ren J, Zhang C. Efficacy of transcranial alternating current stimulation for schizophrenia treatment： A systematic review. J Psychiatr Res. 2023;168：52-63. doi：10.1016/j.jpsychires.2023.10.021

[37] Alexander, M. L., Alagapan, S., Lugo, C. E., Mellin, J. M., Lustenberger, C., Rubinow, D. R., & Fröhlich, F. (2019). Double-blind, randomized pilot clinical trial targeting alpha oscillations with transcranial alternating current stimulation (tACS) for the treatment of major depressive disorder (MDD). Translational psychiatry, 9(1), 106. https：//doi.org/10.1038/s41398-019-0439-0

[38] Grover, S., Nguyen, J.A., Viswanathan, V., & Reinhart, R.M. (2021). High-frequency neuromodulation improves obsessive–compulsive behavior. Nature Medicine, 27, 232 - 238.

[39] Grover S, Fayzullina R, Bullard BM, Levina V, Reinhart RMG. A meta-analysis suggests that tACS improves cognition in healthy, aging, and psychiatric populations. Sci Transl Med. 2023;15(697)：eabo2044. doi：10.1126/scitranslmed.abo2044

[40] Kim J, Kim H, Jeong H, Roh D, Kim DH. tACS as a promising therapeutic option for improving cognitive function in mild cognitive impairment： A direct comparison between tACS and tDCS. J Psychiatr Res. 2021;141：248-256. doi：10.1016/j.jpsychires.2021.07.012

[41] Nissim NR, Pham DVH, Poddar T, Blutt E, Hamilton RH. The impact of gamma transcranial alternating current stimulation (tACS) on cognitive and memory processes in patients with mild cognitive impairment or Alzheimer's disease： A literature review. Brain Stimul. 2023;16(3)：748-755. doi：10.1016/j.brs.2023.04.001

[42] Wang H, Zhang W, Zhao W, et al. The efficacy of transcranial alternating current stimulation for treating post-stroke depression: Study Protocol Clinical Trial (SPIRIT Compliant). Medicine (Baltimore). 2020;99(16): e19671. doi:10.1097/MD.0000000000019671

[43]	Wang HX, Wang L, Zhang WR, et al. Effect of Transcranial Alternating Current Stimulation for the Treatment of Chronic Insomnia: A Randomized, Double-Blind, Parallel-Group, Placebo-Controlled Clinical Trial. Psychother Psychosom. 2020;89(1):38-47. doi:10.1159/000504609

[44]	Antal A, Alekseichuk I, Bikson M, et al. Low intensity transcranial electric stimulation: Safety, ethical, legal regulatory and application guidelines. Clin Neurophysiol. 2017;128(9):1774-1809. doi:10.1016/j.clinph.2017.06.001

__________________________